## Raw Data
Plot Tidal Volumes/Breath vs. Pump Time (y-axis) assuming max-power | [5 empty] | Plot Tidal Volume/sec vs. Pump Power
Initial aperatus weight 790 grams |  |  |  |  | 1 gram of water = 1 milliliter of water
12 breaths per minute at a 1:2 Inhale:exhale ratio | 1 full breath is 5 seconds | 12 breaths per minute at a 1:1.5 Inhale:exhale ratio | inhale for 2 seconds (2000 ms) | 12 breaths per minute at a 1:1.5 Inhale:exhale ratio
Run # | final grams |  | final grams
Run 1 | 173 |  | 208
Run 2 | 165 |  | 210
Run 3 | 179 |  | 210
Run 4 | 177 |  |  |  | volume per breath vs pump run time aka 129mL / 1200ms
Average | 173.5 | Average | 209.3333 |  | bucket weighs 64 grams
Stdev | 6.1914 | Stdev | 1.1547 |  | Glass aparatus weighs 125 grams
14 breaths per minute at 1:2 inhale:exhale ratio | 1 full breath = 60/14 |  | inhale for 1714 ms
Run# | final grams | mL | mL
Run 1 |  |  | 182
Run 2 |  |  | 181
Run 3 |  |  | 183
Average | #DIV/0! | Average | 182
Stdev | #DIV/0! | Stdev | 1
16 breaths per minute at 1:2 inhale:exhale ratio |  |  | 1500ms
Run# | final grams
Run 1 |  |  | 159
Run 2 |  |  | 163
Run 3 |  |  | 155
Run 4 |  |  | 160
Average | #DIV/0! | Average | 159.25
Stdev | #DIV/0! | Stdev | 3.304
18 breaths per minute at 1:2 inhale:exhale ratio
Run# | final grams
Run 1 |  |  | 138
Run 2 |  |  | 136
Run 3 |  |  | 140
Average | #DIV/0! | Average | 138
Stdev | #DIV/0! | Stdev | 2
20 breaths per minute at 1:2 inhale:exhale ratio
Run# |  |  | 129
Run 1 |  |  | 128
Run 2 |  |  | 127
Run 3
Average | #DIV/0! | Average | 128
Stdev | #DIV/0! | Stdev | 1

## Sheet2
Inhale Time (ms) | Average Volume Displaced (mL) | Stdev | Count | Average Flow Rate (L/min) | stdev flow rate
1200 | 128 | 1 | 3 | 6.4 | 0.05
1333 | 138 | 2 | 4 | 6.2116 | 0.090023
1500 | 159.25 | 3.304 | 4 | 6.37 | 0.13216
1667 | 173.5 | 6.1914 | 4 | 6.2448 | 0.22285
1714 | 182 | 1 | 3 | 6.3711 | 0.035006
2000 | 209.3333 | 1.1547 | 3 | 6.28 | 0.034641

## Dawn's Mess
Time(ms) | grams | Volume Displaced (mL) | L/s | L/min | Approx min Pressure (atm) | Approx max Pressure (atm) | Approx min Pressure (Pa) | Approx max Pressure (Pa) | R [(Pa*L)/(K*mol)] | T (F) | T (K) | mol/s (minimum) (PV=nRT; n=PV/RT) | mol/s (maximum) (PV=nRT; n=PV/RT) |  | Double pump mol/s (minimum) | Double pump (L/s) | Double Pump (mL/s) | Seconds for 350mL/breath
1200 | 129 | 129 | 0.1075 | 6.45 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0046248 | 0.0047102 |  | 0.0092496 | 0.215 | 215 | 1.6279
1200 | 128 | 128 | 0.10667 | 6.4 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045889 | 0.0046737 |  | 0.0091779 | 0.21333 | 213.3333 | 1.6406
1200 | 127 | 127 | 0.10583 | 6.35 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045531 | 0.0046372 |  | 0.0091062 | 0.21167 | 211.6667 | 1.6535
1333 | 138 | 138 | 0.10353 | 6.2116 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0044538 | 0.0045361 |  | 0.0089076 | 0.20705 | 207.0518 | 1.6904
1333 | 136 | 136 | 0.10203 | 6.1215 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0043893 | 0.0044703 |  | 0.0087786 | 0.20405 | 204.051 | 1.7153
1333 | 140 | 140 | 0.10503 | 6.3016 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045184 | 0.0046018 |  | 0.0090367 | 0.21005 | 210.0525 | 1.6663
1500 | 159 | 159 | 0.106 | 6.36 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045603 | 0.0046445 |  | 0.0091205 | 0.212 | 212 | 1.6509
1500 | 163 | 163 | 0.10867 | 6.52 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.004675 | 0.0047613 |  | 0.00935 | 0.21733 | 217.3333 | 1.6104
1500 | 155 | 155 | 0.10333 | 6.2 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0044455 | 0.0045276 |  | 0.0088911 | 0.20667 | 206.6667 | 1.6935
1500 | 160 | 160 | 0.10667 | 6.4 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045889 | 0.0046737 |  | 0.0091779 | 0.21333 | 213.3333 | 1.6406
1667 | 173 | 173 | 0.10378 | 6.2268 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0044647 | 0.0045472 |  | 0.0089294 | 0.20756 | 207.5585 | 1.6863
1667 | 165 | 165 | 0.09898 | 5.9388 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0042583 | 0.0043369 |  | 0.0085165 | 0.19796 | 197.9604 | 1.768
1667 | 179 | 179 | 0.10738 | 6.4427 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0046196 | 0.0047049 |  | 0.0092391 | 0.21476 | 214.757 | 1.6297
1667 | 177 | 177 | 0.10618 | 6.3707 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.004568 | 0.0046523 |  | 0.0091359 | 0.21236 | 212.3575 | 1.6482
1714 | 182 | 182 | 0.10618 | 6.3711 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045682 | 0.0046525 |  | 0.0091364 | 0.21237 | 212.3687 | 1.6481
1714 | 181 | 181 | 0.1056 | 6.3361 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045431 | 0.004627 |  | 0.0090862 | 0.2112 | 211.2019 | 1.6572
1714 | 183 | 183 | 0.10677 | 6.4061 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045933 | 0.0046781 |  | 0.0091866 | 0.21354 | 213.5356 | 1.6391
2000 | 208 | 208 | 0.104 | 6.24 | 1.029 | 0.048 | 104263.425 | 4863.6 | 8314.4626 | 65 | 291.4833 | 0.0044742 | 0.00020871 |  | 0.0089484 | 0.208 | 208 | 1.6827
2000 | 210 | 210 | 0.105 | 6.3 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045172 | 0.0046007 |  | 0.0090345 | 0.21 | 210 | 1.6667
2000 | 210 | 210 | 0.105 | 6.3 | 1.029 | 1.048 | 104263.425 | 106188.6 | 8314.4626 | 65 | 291.4833 | 0.0045172 | 0.0046007 |  | 0.0090345 | 0.21 | 210 | 1.6667
[7 empty] | 104.2634 | 106.1886
kPA/100m | kPa/m |  | Pa/m | Pa/ft
116 | 1.16 |  | 1160 | 353.6585
0.049087